# ENDCAP Alignment

6U RS-422 DECOUPLER

##### User’s Manual

FNAL – PNPI

2005

##### Introduction

The RS-422 decoupler provides power-supply decoupling at four identical RS-422 channels. No control is required for this board.

###### Construction

The RS-422 decoupler is 6U 4HP VME module without backplane connectors. Its printed circuit board has size C (233 mm height x 340 mm depth).

The RS-422 decoupler has nine connectors:

J1-J4	– 6-cont. Molex SL headers (RS-422 Input 1 to RS-422 Input 4)

J5	– 4-port RJ12 jack (RS-422 Output 1 to RS-422 Output 4)

J6	– 6-cont. Molex SL header (Low Voltage Input)

- Circuit Diagram

The RS-422 decoupler consists of four identical RS-422 channels, DC/DC converter, and low drop output (LDO) voltage regulator. Each RS-422 channel includes upstream (input) RS-422/CMOS converter MAX3490 (U1 – U4), decoupler ADuM1100 with air core transformer inside (U5 – U12), and downstream (output) CMOS/RS-422 converter MAX3490 (U13 – U16). The DC/DC converter is performed with LTC1474 integrated circuit (U17) and external resistive divider R9, R10 that determines the output voltage as 3.3V. It converts an input 15V voltage to 3.3V to supply downstream circuitry (D3V). The voltage regulator is performed with LT1763 integrated circuit (U18) and external resistive divider R11, R12 that determines the output voltage as 3.3V. It is powered from isolated power supply (F5V, FGND) and supplies upstream circuitry (F3V).

#### Functioning

## The RS-422 decoupler does not need in any control. It works properly if all connections are correct.

#### Debugging and Testing

Necessary equipment:

- Computer running under Windows (9x, NT, 2000)

- RS-232/RS-422 adapter

- DCOPS Interface board #1

- DCOPS Readout board

- 5 – 7 V power supply (F5V)

- 12 – 15 V power supply (D15V)

- RS-232 serial link DB9/DB9 cable

- RS-422 serial link RJ11/Molex SL cable

- RS-422 serial link RJ11/RJ11 cable

- LV cable for RS-232/RS-422 adapter

- LV cable for the interface board

- LV cable for the decoupler under test

- Digital voltmeter

- Oscilloscope

Software to be used:

- Hyper Terminal application

## Properly manufactured RS-422 serial decoupler does not require any adjustments and should work correct at once.

- Check the assembled board visually. It must not have visible damages. All components and their values must be in accordance with documents.

- Connect the decoupler under test to the power supplies and turn on them.

- Check +3.3V output voltage of the DC/DC converter (D3V) in respect to GND.

- Check +3.3V output voltage of the LDO voltage regulator (F3V) in respect to FGND.

- Input nTT command from a keyboard, check a response of the readout board (n = number of the readout board).

- Repeat the test p. 5 for all other inputs/outputs of the decoupler.

Fig. 1. Test of the decoupler. Diagram of connections

PC

RS-232/RS-422

Adapter

DCOPS

Readout Board

DCOPS

Interface

Board # 1

Power Supply

7V

Power Supply

15V

Decoupler

under test

RS-232

IN4

IN3

IN2

IN1

OUT4

OUT3

OUT2

OUT1

LV

- Troubleshooting and Maintenance

  - Output voltage of the DC/DC converter is wrong.

- Check an input voltage; it should be 12 – 15 V.

- Turn off power supplies. Check a resistance between test points D3V and GND, they must not be shortened.

- Check values of the R9 and R10 resistors.

- Check polarity of the C1 and C3 tantalum capacitors (a white band marks a positive pad).

- If the previous checks do not discover a reason of failure, replace the U17 LTC1474 integrated circuit.

  - Output voltage of the voltage regulator is wrong.

- Check an input voltage; it should be 5 – 7 V.

- Turn off power supplies. Check a resistance between test points F3V and FGND, they must not be shortened.

- Check values of the R11 and R12 resistors.

- Check polarity of the C16 and C20 tantalum capacitors (a white band marks a positive pad).

- If the previous checks do not discover a reason of failure, replace the U18 LT1763 integrated circuit.

  - There is no response of the readout board.

- Check correspondence of the command board number to the real setting.

- Check downstream data transfer for the failed channel:

- Press and hold a key “7” at a keyboard,

- Check signal path through:

- input MAX3490 (U1B, U2B, U3B, U4B) pins 8, 7, 2

- ADuM1100 (U6, U8, U10, U12) pins 2, 6

- output MAX3490 (U13A, U14A, U15A, U16A) pins 3, 5, 6.

- Remove a found malfunction.

- Check upstream data transfer for the failed channel:

- Turn off power supplies, change the readout board number to “0” (remove all jumpers) and turn on power supplies again. Now the readout board transfers data continuously.

- Check signal path through:

- input MAX3490 (U13B, U14B, U15B, U16B) pins 8, 7, 2,

- ADuM1100 (U5, U7, U9, U11) pins 2, 6,

- output MAX3490 (U1A, U2A, U3A, U4A) pins 3, 5, 6.

- Remove a found malfunction.

- Specification

General

Dimensions of the board 	13.386" x 9.187"

Input connectors J1-J4	Molex SL header 70553-0005

Output connectors J5		4-port RJ12 jack

Input LV connector J6	Molex SL header 70553-0005

Power characteristics

D15V:

Minimum input voltage			+5 V

Maximum input voltage			+15 V

Output regulated voltage			+3.3 V

Output current consumption:			110 mA

Power consumption:				0.40 W

On board power protection:

LT1474 DC/DC converter has overcurrent and over-temperature protection

F5V:

Minimum input voltage			+5 V

Maximum input voltage			+7 V

Output regulated voltage			+3.3 V

Output current consumption: 			110 mA

Power consumption:				0.77 W

On board power protection:

LT1761 voltage regulator has overcurrent and over-temperature protection

Specification of connectors

Connectors J1 – J4 (IN1 – IN4):

| Pin | Name |
| --- | --- |
| 1 | - |
| 2 | RxD+ |
| 3 | RxD- |
| 4 | TxD+ |
| 5 | TxD- |
| 6 | - |

Connector J5 ports 1 – 4 (OUT1 – OUT4):

| Pin | Name |
| --- | --- |
| 1 | - |
| 2 | RxD+ |
| 3 | RxD- |
| 4 | TxD+ |
| 5 | TxD- |
| 6 | - |

Connector J6 (LV Input):

| Pin | Name |
| --- | --- |
| 1 | D15V |
| 2 | GND |
| 3 | F5V |
| 4 | FGND |
| 5 | - |
| 6 | - |

Fig. 1. 6U RS-422 decoupler