# ENDCAP Alignment

ANALOG READOUT MODULE

##### User’s Manual

FNAL – PNPI

2005

##### Introduction

The analog interface module is intended for measurement of slow changing voltages produced by analog sensors (inclinometer, R, Z, proximity), temperature sensors and for laser control. Thereupon three modification of the module exist:

- 6xAIF		to measure voltages of inclinometer, R-, and Z-sensors

- TC+2PIF 	to measure temperature and voltages of 24 proximity sensors

- 4xPIF 		to measure voltages of 48 proximity sensors

A main measurement/control/interface unit of the module is Embedded Local Monitor Board (ELMB). ELMB has 64 analog inputs and 16 digital outputs (see ELMB User’s Guide). The analog readout module provides matching circuitry for analog inputs and digital outputs controlled by ELMB. The module is supplied with CAN interface and does not require backplane. There are 15 LEDs at the front panel to indicate a state of laser control outputs and presence of proper voltages.

###### Construction

The analog readout module is performed as 6U VME module 340 mm depth. It consists of motherboard, embedded local monitor board (ELMB) inserted into two 100-contact headers, and mezzanine board. The motherboard contains three stacking 40-pin connectors for exchangeable mezzanine board. The mezzanine board provides interconnection between connectors arranged at the front panel and stacking connectors. There are three modifications of the module defined by kind of mezzanine board. The first modification (6xAIF) has six 20-pin headers to connect with analog interface boards. The second modification (TC+2xPIF) has two 26-pin headers to connect with proximity interface boards and six RJ12 jacks for 18 temperature measurement channels. The temperature conversion board is used as the second mezzanine board in this modification. The third modification (4xPIF) has four 26-pin headers to connect with proximity interface boards. The modules 6xAIF and 4xPIF have width 8HP; the TC+2xPIF module has width 12HP.

The analog readout motherboard has eight connectors:

J13, J14	–DB9 CAN bus connectors

J15, J16	– Molex SL 6-pin connectors

J18A, J18B	– 2-port RJ12 jack

Port 1 – Laser Output (This station)

Port 2 – Laser Input

J19, J20	– Molex SL 6-pin connectors

The analog readout board has four switches performed as pin strips controlled by shunts (jumpers):

P1	– Select the input voltage for the Z1 sensor

One shunt is used always

P2	– Connect Z2 input voltage to the output connector J6

Two shunts are used simultaneously. If connection is not required, both

shunts should be placed to a single pin of the same switch.

P3	– Select the current board laser output

Two shunts are used always

P4	– Select a source for the channel 4 at the output connector J6 (pins 7, 8):

Z1 reference voltage or R3 sensor.

Two shunts are used always.

- Circuit Diagram

A main unit of the analog readout module is ELMB inserted into two 100-contact headers. ELMB has 64 analog inputs to measure voltage in a range of 0 to 5VDC. To provide measurement of output voltages for various alignment analog sensors

It consists of the low drop output (LDO) voltage regulator and cross-connection circuitry. Voltage regulator is performed in LT1761 U1 integrated circuit with external resistive divider R1, R2 that determines the output voltage as 11.2 V. Red LED indicates a presence of the regulated voltage.

- Application Notes

- Admissible positions of shunt at the P1 switch are the following (any other position is prohibited):

P1

V1

P1

V2

V2

V1

- Admissible positions of shunts at the P1 switch are the following (any other position is prohibited):

###### P2

###### P2

J6:20

J6:19

V2_GND

J6:20

J6:19

V2

V2_GND

V2

- Admissible positions of shunt at the P1 switch are the following (any other position is prohibited):

###### P2

###### P2

###### P2

L1P

L1M

L2P

L2M

L3P

L3M

L1P

L1M

L2P

L2M

L3P

L3M

L1P

L1M

L2P

L2M

L3P

L3M

- Admissible positions of shunt at the P1 switch are the following (any other position is prohibited):

R3

Z1VREF

R3

Z1VREF

###### P4

###### P4

#### Functioning

## The analog interface board is a passive device representing a cross-connection board with internal voltage regulator. It works properly if all settings and connections are correct.

#### Debugging and Testing

Necessary equipment:

- Digital voltmeter

- 12 – 15 V power supply

- Power cable for analog interface board

## Properly manufactured analog interface boards do not require any adjustments and should work correct at once.

- Check the assembled board visually. It must not have visible damages. All components and their values must be in accordance with documents.

- Connect the analog interface board under test to the power supply 12 – 15 V and turn on it. Check an input voltage after the F1 and F2 fuses; replace them if necessary.

- Check +11.3V output voltage of the LDO voltage regulator.

- A complete test of the analog alignment DAQ including the analog interface board must be performed before a final assembly.

- Troubleshooting and Maintenance

Output voltage of the voltage regulator is wrong.

- Check an input voltage; it should be 12 – 15 V.

- Check a fuse F2; replace it if necessary.

- Check values of the R1 and R2 resistors.

- Check polarity of the C6, C8, and C9 tantalum capacitors (a white band marks a positive pad).

- If the previous checks do not discover a reason of failure, replace the U1 LTC11761 integrated circuit.

- Specification

General

Dimensions of the board 	6.7” x 2.9”

Input connector J1		6-port RJ11 jack

Input/output connector J2	4-port RJ11 jack

Output connector J3		RJ11 jack

Power connectors J4, J5	2-pin Molex Sabre 43160-3102

Output connector J6		20-pin 3M header 3428-6002

Switch P1			3-pin single raw strip 3M 929647-09-36-I

(modified) with shunt 3M 929957-08

Switch P2 			2 x 2 double raw strip 3M 929665-09-36-I

(modified) with shunts 3M 929957-08

Switch P3 			6 x 2 double raw strip 3M 929665-09-36-I

(modified) with shunts 3M 929957-08

Power characteristics

Minimum input voltage			+12 V

Maximum input voltage			+15 V

Output regulated voltage			+11.2 V

Power consumption:				480 mW with entire sensor set

Current consumption:

Inclinometer			       7     mA

R-sensors			2 x 1.13 =   2.26 mA

Z-sensors			2 x 11.3 = 22.6   mA

On board power protection:

Two fuses					0.5 A

LT1761 voltage regulator has overcurrent and overtemperature protection

Specification of connectors

Internal on-board connectors

Stacking connector J1 (channels 1 to 20):

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| 1 | In1+ | 2 | In1- |
| 3 | In2+ | 4 | In2- |
| 5 | In3+ | 6 | In3- |
| 7 | In/Out4+ | 8 | In/Out4- |
| 9 | In5+ | 10 | In5- |
| 11 | In6+ | 12 | In6- |
| 13 | In7+ | 14 | In7- |
| 15 | In8+ | 16 | In8- |
| 17 | In9+ | 18 | In9- |
| 19 | In10+ | 20 | In10- |
| 21 | In11+ | 22 | In11- |
| 23 | In12+ | 24 | In12- |
| 25 | In13+ | 26 | In13- |
| 27 | In14+ | 28 | In14- |
| 29 | In15+ | 30 | In15- |
| 31 | In16+ | 32 | In16- |
| 33 | In17+ | 34 | In17- |
| 35 | In18+ | 36 | In18- |
| 37 | In19+ | 38 | In19- |
| 39 | In20+ | 40 | In20- |

Stacking connector J2 (channels 21 to 40):

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| 1 | In21+ | 2 | In21- |
| 3 | In22+ | 4 | In22- |
| 5 | In23+ | 6 | In23- |
| 7 | In/Out24+ | 8 | In/Out24- |
| 9 | In25+ | 10 | In25- |
| 11 | In26+ | 12 | In26- |
| 13 | In27+ | 14 | In27- |
| 15 | In28+ | 16 | In28- |
| 17 | In29+ | 18 | In29- |
| 19 | In30+ | 20 | In30- |
| 21 | In31+ | 22 | In31- |
| 23 | In32+ | 24 | In32- |
| 25 | In33+ | 26 | In33- |
| 27 | In34+ | 28 | In34- |
| 29 | In35+ | 30 | In35- |
| 31 | In36+ | 32 | In36- |
| 33 | In37+ | 34 | In37- |
| 35 | In38+ | 36 | In38- |
| 37 | In39+ | 38 | In39- |
| 39 | In40+ | 40 | In40- |

Stacking connector J3 (channels 41 to 60):

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| 1 | In41+ | 2 | In41- |
| 3 | In42+ | 4 | In42- |
| 5 | In43+ | 6 | In43- |
| 7 | In/Out44+ | 8 | In/Out44- |
| 9 | In45+ | 10 | In45- |
| 11 | In46+ | 12 | In46- |
| 13 | In47+ | 14 | In47- |
| 15 | In48+ | 16 | In48- |
| 17 | In49+ | 18 | In49- |
| 19 | In50+ | 20 | In50- |
| 21 | In51+ | 22 | In51- |
| 23 | In52+ | 24 | In52- |
| 25 | In53+ | 26 | In53- |
| 27 | In54+ | 28 | In54- |
| 29 | In55+ | 30 | In55- |
| 31 | In56+ | 32 | In56- |
| 33 | In57+ | 34 | In57- |
| 35 | In58+ | 36 | In58- |
| 37 | In59+ | 38 | In59- |
| 39 | In60+ | 40 | In60- |

Pin strip J4 (switch In4   Z1 Laser Control):

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| 1 | In4+ | 2 | In4- |
| 3 | In/Out4+ | 4 | In/Out4- |
| 5 | ZL1Enbl | 6 | GND |

Pin strip J5 (switch In14   Z2 Laser Control):

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| 1 | In14+ | 2 | In14- |
| 3 | In/Out14+ | 4 | In/Out14- |
| 5 | ZL2Enbl | 6 | GND |

Pin strip J6 (switch In24   Z3 Laser Control):

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| 1 | In24+ | 2 | In24- |
| 3 | In/Out24+ | 4 | In/Out24- |
| 5 | ZL3Enbl | 6 | GND |

Pin strip J7 (switch In34   Z4 Laser Control):

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| 1 | In34+ | 2 | In34- |
| 3 | In/Out34+ | 4 | In/Out34- |
| 5 | ZL4Enbl | 6 | GND |

Pin strip J8 (switch In44   Z5 Laser Control):

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| 1 | In44+ | 2 | In44- |
| 3 | In/Out44+ | 4 | In/Out44- |
| 5 | ZL5Enbl | 6 | GND |

Pin strip J9 (switch In54   Z6 Laser Control):

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| 1 | In54+ | 2 | In54- |
| 3 | In/Out54+ | 4 | In/Out54- |
| 5 | ZL6Enbl | 6 | GND |

ELMB-1 connector J10:

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| A1 | - | B1 | - |
| A2 | - | B2 | - |
| A3 | CSA | B3 | - |
| A4 | ADIN | B4 | GND |
| A5 | AMUX | B5 | ACLK |
| A6 | - | B6 | ADOUT |
| A7 | LC2 | B7 | - |
| A8 | LC4 | B8 | LC1 |
| A9 | LC6 | B9 | LC3 |
| A10 | - | B10 | LC5 |
| A11 | - | B11 | - |
| A12 | GND | B12 | GND |
| A13 | ADOUT | B13 | - |
| A14 | AMUX | B14 | - |
| A15 | ADIN | B15 | - |
| A16 | ACLK | B16 | GND |
| A17 | CSA | B17 | - |
| A18 | AGND | B18 | GND |
| A19 | Ch17 | B19 | AGND |
| A20 | Ch18 | B20 | AGND |
| A21 | Ch19 | B21 | AGND |
| A22 | Ch20 | B22 | AGND |
| A23 | Ch21 | B23 | AGND |
| A24 | Ch22 | B24 | AGND |
| A25 | Ch23 | B25 | AGND |
| A26 | Ch24 | B26 | AGND |
| A27 | Ch25 | B27 | AGND |
| A28 | Ch26 | B28 | AGND |
| A29 | Ch27 | B29 | AGND |
| A30 | Ch28 | B30 | AGND |
| A31 | Ch29 | B31 | AGND |
| A32 | Ch30 | B32 | AGND |
| A33 | Ch31 | B33 | AGND |
| A34 | Ch32 | B34 | AGND |
| A35 | Ch49 | B35 | AGND |
| A36 | Ch50 | B36 | AGND |
| A37 | Ch51 | B37 | AGND |
| A38 | Ch52 | B38 | AGND |
| A39 | Ch53 | B39 | AGND |
| A40 | Ch54 | B40 | AGND |
| A41 | Ch55 | B41 | AGND |
| A42 | Ch56 | B42 | AGND |
| A43 | Ch57 | B43 | AGND |
| A44 | Ch58 | B44 | AGND |
| A45 | Ch59 | B45 | AGND |
| A46 | Ch60 | B46 | AGND |
| A47 | AGND | B47 | AGND |
| A48 | AGND | B48 | AGND |
| A49 | AGND | B49 | AGND |
| A50 | AGND | B50 | AGND |

ELMB-2 connector J11:

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| A1 | - | B1 | - |
| A2 | - | B2 | - |
| A3 | - | B3 | - |
| A4 | - | B4 | - |
| A5 | GND | B5 | GND |
| A6 | ZLC1 | B6 | ZLC2 |
| A7 | ZLC3 | B7 | ZLC4 |
| A8 | ZLC5 | B8 | ZLC6 |
| A9 | - | B9 | - |
| A10 | CAN-H | B10 | CAN-L |
| A11 | CAN-GND | B11 | CAN-GND |
| A12 | CAN+7V | B12 | CAN+5V |
| A13 | VREF | B13 | CANRST |
| A14 | D+7V | B14 | D+3.3V |
| A15 | - | B15 | GND |
| A16 | - | B16 | - |
| A17 | A+5V | B17 | A-5V |
| A18 | AGND | B18 | GND |
| A19 | Ch1 | B19 | AGND |
| A20 | Ch2 | B20 | AGND |
| A21 | Ch3 | B21 | AGND |
| A22 | Ch4 | B22 | AGND |
| A23 | Ch5 | B23 | AGND |
| A24 | Ch6 | B24 | AGND |
| A25 | Ch7 | B25 | AGND |
| A26 | Ch8 | B26 | AGND |
| A27 | Ch9 | B27 | AGND |
| A28 | Ch10 | B28 | AGND |
| A29 | Ch11 | B29 | AGND |
| A30 | Ch12 | B30 | AGND |
| A31 | Ch13 | B31 | AGND |
| A32 | Ch14 | B32 | AGND |
| A33 | Ch15 | B33 | AGND |
| A34 | Ch16 | B34 | AGND |
| A35 | Ch33 | B35 | AGND |
| A36 | Ch34 | B36 | AGND |
| A37 | Ch35 | B37 | AGND |
| A38 | Ch36 | B38 | AGND |
| A39 | Ch37 | B39 | AGND |
| A40 | Ch38 | B40 | AGND |
| A41 | Ch39 | B41 | AGND |
| A42 | Ch40 | B42 | AGND |
| A43 | Ch41 | B43 | AGND |
| A44 | Ch42 | B44 | AGND |
| A45 | Ch43 | B45 | AGND |
| A46 | Ch44 | B46 | AGND |
| A47 | Ch45 | B47 | AGND |
| A48 | Ch46 | B48 | AGND |
| A49 | Ch47 | B49 | AGND |
| A50 | Ch48 | B50 | AGND |

Pin strip J12 (Enable CAN bus reset):

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| 1 | Resest In | 2 | Bus reset |

General Front Panel Connectors

CAN bus connectors J13, J14:

| Pin | Name |
| --- | --- |
| 1 | CAN Bus Reset |
| 2 | CAN-L |
| 3 | CAN-GND |
| 4 | - |
| 5 | - |
| 6 | CAN-GND |
| 7 | CAN-H |
| 8 | - |
| 9 | CAN+7V |

Low Voltage connectors J15, J16:

| Pin | Name |
| --- | --- |
| 1 | A+12V |
| 2 | AGND |
| 3 | D+7V |
| 4 | GND |
| 5 | CAN+7V |
| 6 | CAN-GND |

SLM Laser Control connector J18A:

| Pin | Name |
| --- | --- |
| 1 | SLM Laser 1 |
| 2 | GND |
| 3 | SLM Laser 2 |
| 4 | GND |
| 5 | SLM Laser 3 |
| 6 | GND |

SLM Laser Control connector J18B:

| Pin | Name |
| --- | --- |
| 1 | SLM Laser 4 |
| 2 | GND |
| 3 | SLM Laser 5 |
| 4 | GND |
| 5 | SLM Laser 6 |
| 6 | GND |

Transfer Line Laser Control connector J19:

| Pin | Name |
| --- | --- |
| 1 | TL Laser 1 |
| 2 | GND |
| 3 | TL Laser 2 |
| 4 | GND |
| 5 | TL Laser 3 |
| 6 | GND |

Transfer Line Laser Control connector J20:

| Pin | Name |
| --- | --- |
| 1 | TL Laser 4 |
| 2 | GND |
| 3 | TL Laser 5 |
| 4 | GND |
| 5 | TL Laser 6 |
| 6 | GND |

6xAIF Front Panel Connectors

AIF connectors J21 – J26 (AIF1 – AIF6):

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| 1 | Incl VX | 2 | AGND |
| 3 | Incl VY | 4 | AGND |
| 5 | Z1 wiper | 6 | AGND |
| 7 | Z1Vr/R3 | 8 | AGND |
| 9 | Z2 wiper | 10 | AGND |
| 11 | R1 wiper | 12 | AGND |
| 13 | R1 Vref | 14 | AGND |
| 15 | R2 wiper | 16 | AGND |
| 17 | R2 Vref | 18 | AGND |
| 19 | Z2 Vref | 20 | AGND |

TC+2xPIF Front Panel Connectors

PIF connectors J28, J29 (PIF1, PIF2):

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| 1 | VPX1-1 | 2 | AGND |
| 3 | VPX1-2 | 4 | AGND |
| 5 | VPX2-1 | 6 | AGND |
| 7 | VPX2-2 | 8 | AGND |
| 9 | VPX3-1 | 10 | AGND |
| 11 | VPX3-2 | 12 | AGND |
| 13 | VPX4-1 | 14 | AGND |
| 15 | VPX4-2 | 16 | AGND |
| 17 | VPX5-1 | 18 | AGND |
| 19 | VPX5-2 | 20 | AGND |
| 21 | VPX6-1 | 22 | AGND |
| 23 | VPX6-2 | 24 | AGND |
| 25 | Vref | 26 | AGND |

Temperature conversion ports 1 to 6:

| Pin | Name |
| --- | --- |
| 1 | T1 |
| 2 | Vref |
| 3 | T2 |
| 4 | Vref |
| 5 | T3 |
| 6 | Vref |

4xPIF Front Panel Connectors

PIF connectors J30 – J33 (PIF3 – PIF6):

| Pin | Name | Pin | Name |
| --- | --- | --- | --- |
| 1 | VPX1-1 | 2 | AGND |
| 3 | VPX1-2 | 4 | AGND |
| 5 | VPX2-1 | 6 | AGND |
| 7 | VPX2-2 | 8 | AGND |
| 9 | VPX3-1 | 10 | AGND |
| 11 | VPX3-2 | 12 | AGND |
| 13 | VPX4-1 | 14 | AGND |
| 15 | VPX4-2 | 16 | AGND |
| 17 | VPX5-1 | 18 | AGND |
| 19 | VPX5-2 | 20 | AGND |
| 21 | VPX6-1 | 22 | AGND |
| 23 | VPX6-2 | 24 | AGND |
| 25 | Vref | 26 | AGND |

Fig. 1. The analog readout motherboard

Fig. 2. The analog readout module 6xAIF

Fig. 3. The analog readout module TC+2xPIF

Fig. 4. The analog readout module 4xPIF